Change Control SOP 016827 REV CC - Accuray Confidential Copyright 2025 Accuray Incorporated

Revision History

| Rev | Change # | Description of Change | Revised By | Date |
| --- | --- | --- | --- | --- |
| CA | DCO-01397 | Move definitions to 1089757, Change "Product Development" section to "Change Control Management" and move relevant data to 1090230, QMS Change Process removed, “Release of Items (CCB)” and “Implementation of Items” moved to 1052622, remove references to change type DCO. | Nathan Shirey | 05-Jul-2023 |
| CB | CHG-00222 | Revise Design Change Impact Risk process. Product Change Process moved to WI 1063933. Clarification of process for non-project changes. Removed unused reference documents from Section 3. | Nathan Shirey | 01-May-2024 |
| CC | CHG-01171 | Remove references to Tier 3s, move actions for “Product Change Process ECO approvers” from 1063933 into 016827. Product Change Process detailed moved into 016827 from 1063933. | Nathan Shirey | See DMS |

Table of Contents

1	Purpose	4

2	Scope	4

3	Reference Documents	5

4	Definitions	5

5	Procedure	6

5.1	Change Control Processes	6

5.2	Change Control Process Types of Changes	8

6	Records	12

7	Appendix	13

7.1	Appendix A – Administrative Changes	13

# Purpose

The purpose of this procedure is to define Accuray’s process for the initiation and approval of design and lifecycle changes relating to component, material, product, assembly, process, or quality system to ensure that the impact of these changes on product safety, efficacy, or quality has been formally reviewed and authorized prior to implementation.

# Scope

  - This procedure applies to finished design outputs of Accuray products, bills-of-material, processes, components, and equipment as well as related documentation and Quality System documentation (see Figure 1: Change Control and Process Governance).

  - This procedure applies to all departments handling activities as defined in Section 2.1.

  - The procedures described in this document reflect the standard process by which items are created and changed within the function of Change Control.  The following change procedure is not included within the scope of this document and is within another function.

- SOP 026921 – IT

  - This procedure does not apply to company collateral and promotional materials.

Figure 1: Change Control and Process Governance

# Reference Documents

- SOP 016087 Item Numbering and Naming

- SOP 016824 Quality Records and Records Retention

- SOP 020761 Deviation

- SOP 023621 Product Hold

- SOP 026921 CSV Change Control

- SOP 1004376 Document Control

- SOP 1049123 Product Development Process

- WI 018949 Design History File

- WI 032394 Creating ECO Items and Changes

- WI 1040306 Agile Employee User Guide

- WI 1041758 Third Party Access to Agile

- WI 1046343 Creating and Managing ECRs

- WI 1063718 QMS Documents

- WI 1063933 Product Change Review Board

- WI 1063971 Using SolidWorks with Agile

- WI 1066466 Creating DCRs

- WI 1066467 WI Creating MasterControl Documents and Changes

- WI 1069247 FRU Development and Release Process

- FRM 033017 Initial Change Assessment

- TMPLT 1070627 Field Implementation Plan

- RECORD 1052622 Agile Change Reviewer Matrices

- EU GMP Annex 15: Qualification and Validation

- 21 CFR 820

# Definitions

See Accuray Acronyms and Definitions (1089757).

# Procedure

## Change Control Processes

All changes applicable to this procedure [including “Prototype” lifecycle and Product Configurator Tool (Product Builder)] shall additionally follow (at a minimum, where applicable) Sections 5.2-5.3 for creation, modification, approval, release, and implementation of Change Orders.

### Change Management

#### The PDP (governing procedure) uses Change Control as the tactical procedure for defining and managing changes for parts, process, and product related documentation for product over phases 1-7 of the product lifecycle.  Change Control is both a governing and tactical procedure for design change control from P7 and beyond (project related changes), although there is overlap during P7 with PDP governance.

#### Non-project Changes do not require a Project Charter or Project Initiation checklist. This type of change shall follow WI 1046343 Creating and Managing ECRs and WI 032394 Creating ECO Items and Changes.

#### Prior to full approval of the respective ECR(s), ECO(s) may be initiated and populated. Any information populated in the ECO(s) must be high enough level that it would not be impacted by decisions or changes which would result from the CRB.

### Classification of Design Change Impact Risk

#### Risk based impact determined for all applicable proposed changes with a risk level impact of High or Low.

#### Proposed changes will be evaluated for High or Low impact to Product and Process with consideration for impact to Product Safety Risk, Regulatory Compliance Risk, and Product Quality. This evaluation may be done using FRM 033017. Appropriate actions to mitigate any increased level of risk shall be taken.

### Product Change Process ECO approvers shall:

#### Verify ICA content and Change Owner (or designee) Design Change Impact Risk – Low impact determination is accurate (if applicable).

#### Verify actions defined by CRB approved ICA are addressed with appropriate objective evidence to demonstrate conformance for all required change actions.

#### Identify any approved ICA plan actions that may have been altered or omitted in execution evidence and change order documentation from the CRB ICA approval submission of the ECO by the Change Owner (or designee) and verify there is appropriate documented rationale commensurate with the Design Change Impact Level of High determined by CRB, or the change must be rejected (minimum of one approver) and brought back to CRB for further review and determination/alteration of the approved ICA.

#### Identify changes in scope affecting the approved ICA and verify if a change has been altered in a manner that could increase the CRB design change impact risk level determination or driving factors, the change must be rejected (minimum of one approver) and brought back to CRB for further review and determination/alteration of the approved ICA.

#### Confirm that the Related Project ID is used for all ECOs for designated projects.

## Change Control Process Types of Changes

### Types of changes covered under the scope of this procedure:

- ECO – Engineering Change Order

- ECO-DHF – Engineering Change Order-Design History File

- CHG – Change Packet

- MCO – Manufacturing Change Order

- MCO-DHF – Manufacturing Change Order-Design History File

- ECR – Engineering Change Request

- DCR – Document Change Request

### Types of changes not covered under the scope of this procedure:

- DEV – Deviation (see SOP 020761)

- PH – Product Hold (see SOP 023621)

### General Requirements

#### All changes are created and approved via the appropriate document management system [or Product Lifecycle Management (PLM) software].

#### The Originator/Initiator is responsible for completing all fields of the change, attachments, and submission.

#### The following are responsible for review of the change in accordance with this SOP, SOP 1004376, and related WI.  Reviewers will verify the completeness of the change and notify the Originator/Initiator of any corrections required.  Once complete, reviewers will move the change into the approvals step of the workflow for Technical Review.

- Document Control for CO (WI 032394), CO-DHF (WI 018949), DCR (WI 1066466), and QMS CHG (WI 1066467 and WI 1063718)

- Designee for ECR (WI 1046343)

- Initiator for non-QMS CHG (WI 1066467)

#### Technical Review and approval of the change is reviewed within the document management system by a Change Control Board (CCB) or required reviewers.  The required reviewers are defined in RECORD 1052622.

#### If a change is rejected by a reviewer, it will be returned to the Originator/Initiator until they are prepared to resubmit with any required revisions.

#### Once a change is approved, it will be Released.

#### For changes in the approvals state, where minor errors are discovered or supporting documentation is requested during review, Document Control may (or may not at Document Control’s discretion or DMS functionality) correct or add supporting (Attachments) documentation where corrections/additions do not change the scope or intent of the change (additions) and any documentation (corrections).  Document Control shall confirm with QA prior to making the changes if they need additional guidance and provide objective evidence to reviewers detailing all changes with rationale, where appropriate.

#### After a change is Released, there may be requisite actions for implementation.  Once finished, the change is complete.

#### THAP or Test House Audited Parts: Selected parts are listed in registration documents, compliance filing, and certifications. These parts are identified in the PLM system so that the associated regulatory and compliance reports are reviewed and an assessment executed to determine what actions are needed (e.g., testing, re-registration, document update, etc.).

### ECOs – Results in a new or revised part or product document.  ECOs are required for:

- Alpha or Numeric Revision Changes

- F/F/F Changes

- Item Attachment Changes

- BOM Changes

- Custom Manufacturer Parts

NOTE: Changes to First Article Inspection (FAI) and stock dispositions in the ECO Affected Items tab can only occur in an ECO when changing the actual item and cannot be changed solely on their own.  For non-revision items that are affected, FAI and Stock disposition impact shall be documented on the ECO Cover Page (i.e., items redlined off of a BOM that would require rework to a new/different PN and revision for Inventory and Production/Service use).

See WI 032394 for additional requirements.

### ECO-DHF – Results in a new or revised DHF document.

See WI 018949 for additional requirements.

### MCOs – Results in a change to a released part or product document without changing the revision, item’s Attachments tab, BOM, or FAIR.  MCOs are only to be used to make changes that do not affect F/F/F.

#### MCOs are used to:

- Change Lifecycle Phase to Obsolete, Inactive, or from Pre-Production to Production

- PWBs only: Change Lifecycle Phase from Prototype to Pre-Production or Production

- Change OTS Manufacturer part, where appropriate design verification evidence has been provided (e.g., replacing a Manufacturer part with an equivalent Manufacturer part, adding alternate Manufacturer parts to an item, removing Manufacturer parts, adding RoHS certificates, etc.)

- Change item attributes in the Title Block tab

- Change the RoHS Compliant field to Yes if compliance can be demonstrated with and evidence provided for current revision drawings/specifications

#### MCOs may not be used to (ECO must be used):

- Change Item Attachments Tab (MCO cannot modify)

- Change Custom Manufacturer Part [Modify/Replace a PN, manufacturer/supplier (including alternates)]

#### See WI 032394 for additional requirements.

### MCO-DHFs – Results in a change to a released DHF document without changing the revision, document’s Attachments tab, or BOM (if applicable).

#### MCO-DHFs are used to:

- Change Lifecycle Phase to Obsolete

- Change item attributes in the Title Block tab

#### MCO-DHFs may not be used to:

- Change document Attachments Tab (MCO-DHF cannot modify)

#### See WI 018949 for additional requirements.

### ECRs – Documents a change request, change review, or proposal for a part or product document, including Product Change Review Board (PCRB) review of parts or product documents.  ECRs are used to review the concept of proposed changes, but they cannot make changes.

#### ECRs are required as defined per WI 1046343.

#### ECRs will be reviewed in accordance with WI 1046343.

### DCRs – Documents an administrative change request for parts, product documents, DHF documents, non-product documents, or QMS documents.  Documents a change request, change review, or proposal for QMS CRB review of QMS Quality Manual, SOPs, and WIs.  DCRs are used to review the concept of proposed changes, but they cannot make changes.

#### DCRs are required for:

- Revisions to the Quality Manual

- Initial release, revision, or obsolescence of QMS SOPs/WIs

- Administrative Changes

#### Document Control will review submitted QMS CRB DCRs for scope.  If the DCR reflects small corrections that do not change the scope or intent of a document, Document Control will make the change as an Administrative Change.  See Appendix A – Administrative Changes.

#### QMS CRB DCRs for the Quality Manual and QMS SOPs/WIs will be reviewed by the QMS Change Review Board (CRB) as outlined in WI 1063718.

#### DCRs are not required to initiate a change for all other QMS documents (i.e., Forms, Templates, Quality Policy, and Quality Plans).

#### See WI 1066466 for additional requirements.

### The following documents have 016827 as the parent SOP:

- SOP 016087 Item Numbering and Naming

- WI 032394 Creating ECO Items and Changes

- WI 1040306 Agile Employee User Guide

- WI 1041758 Third Party Access to Agile

- WI 1046343 Creating and Managing ECRs

- WI 1063933 Product Change Review Board

- WI 1063971 Using SolidWorks with Agile

- WI 1066466 Creating DCRs

- WI 1069247 FRU Development and Release Process

- TMPLT 1070627 Field Implementation Plan

- FRM 033017 Initial Change Assessment

# Records

Refer to SOP 016824.

# Appendix

## Appendix A – Administrative Changes

Administrative changes may be used for proposed changes on a DCR that do not change the scope or intent of a document and pose no risk.  Administrative changes must include justification.  The following are examples of changes that can be made administratively through DCR.

- Spelling, typos, grammar, punctuation, formatting, or identification data corrections that do not affect intent

- Change to template attributes (e.g., company logo)

- Change of Department Name (only if the Department change does not affect the roles or responsibilities of the Department)

- Update to Reference Document Titles or Part Numbers (where applicable) (to align with Documents that are re-named or require new Part Numbers)

- Removal of obsolete documents (when not imbedded in the content of the document)