Change is potentially significant for a region and requires review by local RA SME

Checklist requires Regional RA review?

Process Flow

Product RA Evaluate Change using Checklist

Yes

No

Change has no RA impact. Sign off on change

Change Reference Number:

Impacted Product Name/Family:

Instructions: Complete the checklist following the instructions in the ‘Conditions’ column. Note: If the decision to any question is to route for Regional RA assessment, the rest of the checklist doesn’t need to be filled in. Instead, complete the conclusion and initiate Regional RA assessments.

This checklist was based on the following:

- FDA Guidance on Deciding When to Submit a 510(k) for a Change to an Existing Device issued October 25, 2017

- FDA Guidance on Deciding When to Submit a  510(k) for a Software Change to an Existing Device issued October 25, 2017

- MDCG 2020-3 Rev.1  Guidance on significant changes regarding  the transitional provision under Article 120  of the MDR with regard to devices covered  by certificates according to MDD or AIMDD issued May 2023

- RDC 340/2020 and IN 61/2020,Brazil

- PMDA??

- Changes affecting TGA-issued conformity assessment certificates Guidelines for notifying the TGA about ‘substantial changes’ to, or transfers of, conformity assessment certificates Version 2.0 issued August 2021

- Health Canada Guidance for the Interpretation of Significant Change of a Medical Device issued January 20, 2011

- CHINA (??)

- SINGAPORE (??)

| Regulatory Impact Assessment Pre-Screen Checklist |  |  |  |
| --- | --- | --- | --- |
| Question | Conditions | Yes/No Decision | Comments (if applicable) |
| - Is this change due to a recall? | Yes – Route to Regional RA for assessment No – Answer #2 |  |  |
| - Change made with intent to significantly improve the safety or effectiveness of the device? | Yes – Route to Regional RA for assessment No – Answer #3 |  |  |
| - Labeling change? | Yes – Answer #3a No – Answer #4 |  |  |
| - Is it a change in the intended use/purpose or indications for use statement, which is not clerical in nature? | Yes – Route to Regional RA for assessment No – Answer #3b |  |  |
| - Does the change add or delete a contraindication? | Yes – Route to Regional RA for assessment No – Answer #3c |  |  |
| - Is it an addition, deletion or change in the warnings or precautions, which is not clerical in nature? | Yes – Route to Regional RA for assessment No – Answer #3d |  |  |
| - Is there a change to the patient population, a new disease condition, or new way of clinical application? | Yes – Route to Regional RA for assessment No – Answer #3e |  |  |
| - Could the change affect the directions for use? |  |  |  |
| - Is there a change from a device labeled for single use to a device labeled as reusable? |  |  |  |
| - Is it a change from Rx to over the counter OTC use? |  |  |  |
| - Does the change impact the product name, model number, or part number labeled on the device or referenced in labeling? | Yes – Route to Regional RA for assessment No – Answer #3f |  |  |
| - Does the change result in an address change of the Legal Manufacturer and/or the Authorized Rep on the labeling? | Yes – Route to only EU Regional RA for assessment and continue to answer #3g No – Answer #3g |  |  |
| - Does the change affect any labeling (e.g. IFU, PEG, User Interface, product labels etc) contents intended for product configurations marketed in China? | Yes – Route to only China Regional RA for assessment and continue to answer #4 No –Answer #4 |  |  |
| - Technology, engineering or performance change? | Yes – Answer #4a No – Answer #6 |  |  |
| - Is it a control mechanism, operating principle, alarms, or energy type change? | Yes – Route to Regional RA for assessment No – Answer #4b |  |  |
| - Is it a change in sterilization, cleaning, or disinfection? | Yes – Answer #4.b.i No – Answer #4.c |  |  |
| - Is it a change to a Category B or novel method, does it lower the SAL, or is it a change to how the device is provided? | Yes – Route to Regional RA for assessment No – Answer #4.b.ii |  |  |
| - Could the change significantly affect performance/ biocompatibility? | Yes – Route to Regional RA for assessment No – Answer #4.b.iii |  |  |
| - Is there a change that increases the bioburden alert or action levels or that introduces a more difficult to kill organism? | Yes – Route to Regional RA for assessment No – Answer #4.b.iv |  |  |
| - Is there a device design or material change that introduces a more difficult to sterilize feature? | Yes – Route to Regional RA for assessment No – Answer #4.b.v |  |  |
| - Is there a change in sterilization process (e.g. cycle parameters, equipment, density or configuration of load)? | Yes – Route to Regional RA for assessment No – Answer #4.b.vi |  |  |
| - Is a change made from a non-parametric release to a parametric release? | Yes – Route to Regional RA for assessment No – Answer #4.c |  |  |
| - Is there a significant change in packaging or change to expiration dating? | Yes – Route to Regional RA for assessment No – Answer #4.d |  |  |
| - Is there a change to the storage, transportation conditions or shelf-life of the device? | Yes – Route to Regional RA for assessment No – Answer #4.e |  |  |
| - Is there any other change in design (e.g., dimensions, performance specifications, wireless communications, components or accessories, user interface, etc)? | Yes – Answer #4.e.i No – Answer #4.f |  |  |
| - Does the change significantly affect the use of the device? | Yes – Route to Regional RA for assessment No – Answer #4.e.ii |  |  |
| - Does a risk assessment identify any new or significantly modified risks? | Yes – Route to Regional RA for assessment No – Answer #4.e.iii |  |  |
| - Is clinical data necessary? | Yes – Route to Regional RA for assessment No - Answer #4.e.iv |  |  |
| - Any unexpected issues from Verification & Validation activities? | Yes – Route to Regional RA for assessment No – Answer #4.f |  |  |
| - Is clinical data necessary to support the safety and effectiveness of the altered device? | Yes – Route to Regional RA for assessment No – Answer #4.g |  |  |
| - Do the results of a risk analysis, undertaken during the design validation process, raise new issues of safety and effectiveness? | Yes – Route to Regional RA for assessment No – Answer #4.h |  |  |
| - Does the change affect the current PTR and is not clerical/administrative in nature? | Yes – If China is an impacted market of the change, route only to China Regional RA for assessment and continue to answer #5 No – Answer #5 |  |  |
| - Materials or Substance change? | Yes – Answer #5a No – Answer #6 |  |  |
| - Does a risk assessment identify any new or increased biocompatibility concerns? | Yes – Route to Regional RA for assessment No – Answer #5b |  |  |
| - Could the change affect performance specifications or material specifications? | Yes – Route to Regional RA for assessment No – go to #5c |  |  |
| - Change to a material of human/animal origin or a medicinal substance, including addition of new materials? | Yes – Route to Regional RA for assessment No – Answer #5d |  |  |
| - Change to the excipient of a medicinal substance or to the medicinal substance itself? | Yes – Route to Regional RA for assessment No – Answer #5e |  |  |
| - Change to a surgically invasive device’s material or substance in contact >30d with patient tissue or flued or absorbed? | Yes – Route to Regional RA for assessment No – Answer #5f |  |  |
| - Change affects the safety or performance and negatively affects the benefit/risk ratio of the device? | Yes – Route to Regional RA for assessment No – Answer #6 |  |  |
| - Software change? | Yes – Answer #6a No – Answer #7 |  |  |
| - New or major change of operating system or any critical component (e.g. THAP or component listed on China PTR)? | Yes – Route to Regional RA for assessment No – Answer #6b |  |  |
| - New or major modification of architecture, database structure, or change of core functionality (i.e. algorithm)? | Yes – Route to Regional RA for assessment No – Answer #6c |  |  |
| - Required user input replaced by closed loop algorithm (e.g. fully automated)? | Yes – Route to Regional RA for assessment No – Answer #6d |  |  |
| - New medical, diagnostic, or therapeutic feature, removal of alarm, or new channel of inter-operability? | Yes – Route to Regional RA for assessment No – Answer #6e |  |  |
| - New user interface or presentation of data (i.e. medical data represented in new format or by a new dimension/measuring unit)? | Yes – Route to Regional RA for assessment No – Answer #6f |  |  |
| - Does the change introduce a new risk or modify an existing risk that could result in significant harm and that is not effectively mitigated in the most recently cleared device? Or Does the change create or necessitate a new risk control measure or a modification of an existing risk control measure for a hazardous situation that could result in significant harm? | Yes – Route to Regional RA for assessment No – Answer #6g |  |  |
| - Could the change significantly affect clinical functionality or performance specifications that are directly associated with the intended use of the device? | Yes – Route to Regional RA for assessment No – Answer #6h |  |  |
| - Is there a change to programming language (e.g. C to C++) or involve significant software code re-write (e.g. re-engineering)? | Yes – Route to Regional RA for assessment No – Answer #6i |  |  |
| - Is there any update to the major software release version? | Yes – Route to Regional RA for assessment No – Answer #7j |  |  |
| - Is there any update to the minor software release version? | Yes – XXX No – Answer #8 |  |  |
| - Manufacturing Process, Facility, or Equipment Change? | Yes – Answer #7a No – Answer #8 |  |  |
| - Is there a change to a viral inactivation process? | Yes – Route to Regional RA for assessment No – Answer #7b |  |  |
| - Is there a new manufacturing facility being added for the device? Note: This includes new or moved sites that requires address changes, as well as new sites where the address on the label remains the same, but a new mfg facility is added. | Yes – Route to Regional RA for assessment No – Answer #7c |  |  |
| - Has there been a change to incoming inspections to evaluate the material/equipment from a supplier? | Yes – Route to Regional RA for assessment No – Answer #7d |  |  |
| - Is the change being made to a combination product? | Yes – Route to Regional RA for assessment No – Answer #7e |  |  |
| - Is there a change to a drug coating process? | Yes – Route to Regional RA for assessment No – Answer #7f |  |  |
| - Is there a change in the packaging process that impacts the sterile integrity of the device? | Yes – Route to Regional RA for assessment No – Answer #7g |  |  |
| - Is there a change to a manufacturing process that may impact pre-sterilization bioburden? | Yes – Route to Regional RA for assessment No – Answer #8 |  |  |
| - Manufacturing Quality Control Procedure Change? | Yes – Answer #8a No – Checklist completed. There is no regulatory impact for any regions. |  |  |
| - Is the change due to removal of test acceptance criteria, in-process inspections, or final inspections without replacement of these activities? | Yes – Route to Regional RA for assessment No – Checklist completed. There is no regulatory impact for any regions. |  |  |
| - Sterilization or Packaging Change that impacts sterilization? |  |  |  |
| - Is there a change to the terminal sterilization method? |  |  |  |
| - Could the change adversely affect the sterility assurance levels? |  |  |  |
| - Change in packaging desigh, which affects sterility, stability or microbiloical state of the device, including seal integrity |  |  |  |
| - Shelf-life extension? |  |  |  |
| Conclusion: Check one box below. ☐ No regulatory impact for any regions. |  |  |  |
| ☐ Regional RA Assessment required for only one or more of the following regions: China, EU xx. No impact in any other regions. |  |  |  |
| ☐ Regional RA Assessment required for all regions. |  |  |  |
| Approval |  |  |  |
| TITLE PRINT NAME SIGNATURE DATE |  |  |  |